1.3   Technical Specifications

| PURPOSE OF USE |  |  |
| --- | --- | --- |
| 1 | Clinical or other purpose | The cobas 6800 system automated and integrated workflow to run Polymerase Chain Reaction (PCR) based Nucleic Acid Testing (NAT) to be applied in diagnostic and blood screening laboratories. The Systems combines Instrumentation, Consumables, Reagents and Data Management to provide an efficient workflow from sample processing to result interpretation. |
| 2 | Level of use | Health centers, district hospitals, specialized hospitals, laboratories, transfusion centers. |
| 3 | Overview of functional requirements | The cobas® 6800 system automate the preparation and analysis of samples for quantitative and qualitative nucleic acid testing using real-time PCR technology. The cobas® 6800/8800 Systems combine the functionalities of instrumentation, consumables, reagents and data management to provide an efficient workflow from sample processing to result interpretation. The cobas® 6800 Systems and cobas® 8800 Systems are to be used by laboratory professionals trained in laboratory techniques and by instruction on the use of the instrument. |

| TECHNICAL CHARACTERISTICS |  |  |
| --- | --- | --- |
| 4 | Detailed requirements | The cobas® 6800/8800 Systems are designed to run Polymerase Chain Reaction (PCR) based Nucleic Acid Testing (NAT) to be applied in diagnostic and blood screening laboratories. Delivers the fastest time to results, completing up to 96 tests in less than 3.5 hours, maximum throughput (8 hrs ¦ 24 hrs) - 384 ¦ 1,344 tests. Maximum number of tests per run: 96 tests including controls. |
| 5 | Displayed parameters | Digital display of test results. |
| 6 | User adjustable settings | Cobas 6800 completely automatic “close” system without User interaction. |
| PHYSICAL/CHEMICAL CHARACTERISTICS |  |  |
| 7 | Components | Each System is comprised of: - Sample Supply Module - which loads and unloads sample tubes to and from the Transfer Module - Transfer Module - which pipettes samples and controls from the sample and control tubes in processing plates - Processing Modules - which perform sample preparation and nucleic acid extraction - Analytic Modules - which perform real-time PCR on the processed samples - The cobas® 6800 system also include the system software, cobas p 680 instrument (optional for Blood screening) and the Instrument Gateway Server |
| 8 | Mobility, portability | Tabletop. cobas 6800 can be fixed or movable |
| UTILITY REQUIREMENTS |  |  |
| 10 | Electrical, water and/or gas supply (if relevant) | Line voltage: 200-240 VAC +/- 10% Line frequency: 50/60Hz +/- 5% Insulation coordination: Installation category II (IEC 61010-1) Electrical protection provided by fuses in both live and neutral supply lines. Mains cable to be at least 3m in length. Protections against over-voltage and over-current line conditions. Compliance with applicable ____ standards and regulations. Equipment provided with an adequate Uninterruptible Power Supply (UPS) system with not less than 20 minutes of back-up. |
| ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS |  |  |
| 11 | Accessories (if relevant) |  |
| 12 | Sterilization process for accessories (if relevant) |  |
| 13 | Consumables / reagents (if relevant) | Transforms “walk-away” time into true “work-away” time by minimizing user interactions: - Ready-to-use reagents do not require thawing, mixing or pouring - Automated onboard storage and refrigeration system enable ready - access and maintain inventory of consumables and reagents - Radio-frequency identification (RFID) and barcodes ensure full traceability from sample in to results out - Uni- and bi-directional LIS interface simplifies order and result handling |
| 14 | Spare parts (if relevant) |  |
| 15 | Other components (if relevant) |  |
| PACKAGING |  |  |
| 16 | Sterility status on delivery (if relevant) | N/A |
| 17 | Shelf life (if relevant) | Min. 6 month |
| 18 | Transportation and storage (if relevant) | Approx. +2C to +8C degree |
| 19 | Labelling (if relevant) | N/A |
| ENVIRONMENTAL REQUIREMENTS |  |  |
| 20 | Context-dependent requirements | Ambient room temperature: 15°C to 28°C Relative humidity: 30% to 80% (no condensation) Air pressure: 80-106 kPa Altitude: <2000m |
| TRAINING, INSTALLATION AND UTILISATION |  |  |
| 21 | Pre-installation requirements | Local clinical staff to affirm completion of installation |
| 22 | Requirements for commissioning | Supplier to perform installation (drainage), safety and operation checks before handover Local clinical staff to affirm completion of installation |
| 23 | Training of user/s | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented |
| 24 | User care | Casing to be cleanable with alcohol and chlorine wipes |
| WARRANTY AND MAINTENANCE |  |  |
| 25 | Warranty | It depends on agreement terms |
| 26 | Maintenance tasks | Daily, weekly and yearly maintenance are performed. No need calibration. |
| 27 | Type of service contract | It depends on agreement terms |
| 28 | Spare parts availability post-warranty | It depends on agreement terms |
| 29 | Software / Hardware upgrade availability | It depends on agreement terms |
| DOCUMENTATION |  |  |
| 30 | Documentation requirements | User, technical and maintenance manuals to be supplied in local language. List to be provided of equipment and procedures required for local and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. |
| DECOMMISSIONING |  |  |
| 31 | Estimated Life Span | At least 5 years |
| SAFETY AND STANDARDS |  |  |
| 32 | Risk Classification | Principles and Guidelines, International Organization for Standardization (ISO 31000, 2009) |
| 33 | Regulatory Approval / Certification | The cobas 6800 system and the reagents have CE-IVD certificates showing that they are for diagnostic use as per the 98/79/EC Directive on In Vitro Diagnostic Medical Devices, and it is documented that they have TİTUBB/UTS registration |
| 34 | International standards | EN ISO 13485:2016 Medical devices – Quality management systems – Requirements for regulatory perposes ISO 14001:2015 Environmental Management system (will be expired 2020-06-30, new certificate will be available) |
| 35 | Regional / Local Standards |  |
| 36 | Regulations |  |