## Analyzers
Point of Care Blood-Gas Analyzers
|  | Handheld | Compact (tabletop) | Single/Multi Cartridge | Key Measured Parameters | Accepted blood sources | Time to Result | Required sample size | Calibration process | Quality control | Battery-powered | Integrated bar code scanner | Integrated Printer | Database integration | Warranty (months)
Siemens | ePOC | ✔ |  | Single sample | pH, pCO2, pO2, TCO2 electrolytes Hemocrit Glucose, Lactate, Urea | Whole blood: arterial, venous, mixed venous, capillary | < 1 minute | > 100 μl | automatically between tests | Eurotrol GAS-ISE Metabolite control Eurotrol Hematrocrit control | Yes - 50 test cards | Yes | Bluetooth printer accessory | WiFi | 12
Siemens | RapidPoint 500e |  | ✔ | Multiple sample variable sizes: 100/250/400/750 tests 28 day useful life | pH, pCO2, pO2, electrolytes Glucose, Lactate, tHb, CO-oximetry, nBili | Whole blood: arterial, venous, mixed venous, capillary pleural fluid dialysate | 1 minute | > 100 μl | 1-point every 30 minutes 2-point every 2 hours full calibration every 8 hours | Programmable and customizable automatic QC | Wall-powered | Yes | Yes | WiFi, ethernet, serial port | 12
Auto Hematology Analyzers - Complete Blood Count
|  | Size | Number of Measured Parameters | Key Measured Parameters | 3-part or 5-part CBC Differential | Interfaces/Communication | Throughput Samples/hour | Required sample size | Data Storage | Reagants needed | Printout | Warranty (months)
Mindray | BC 10 | Tabletop | 20 3 histograms | WBC, RBC, HGB, MCV, PLT | 3-part | 4 USB ports, WiFi, LAN port | 30 | Prediluated: ≤ 20 μl Whole blood: ≤ 9 μl | 50,000 results | 3 total: M-30D Diluent M10CFL Lyse Probe cleanser | Integrated thermal printer | 18
Mindray | BC 20s | Tabletop | 19 3 histograms | WBC, RBC, HGB, MCV, PLT | 3-part | 4 USB ports, WiFi, LAN port | 40 | Prediluated: ≤ 20 μl Whole blood: ≤ 9 μl | 200,000 results | 3 total: M-30D Diluent M30CFL Lyse Probe cleanser | Integrated thermal printer | 18
Mindray | BC 30s | Tabletop | 21 3 histograms | WBC, RBC, HGB, MCV, PLT | 3-part | 4 USB ports, WiFi, LAN port | 70 | Prediluated: ≤ 20 μl Whole blood: ≤ 9 μl | 500,000 results | 3 total: M-30D Diluent M30CFL Lyse Probe cleanser | Integrated thermal printer | 18
Mindray | BC 2800 | Tabletop | 19 3 histograms | WBC, RBC, HGB, MCV, PLT | 3-part | 2 RS232 | 30 | Prediluted: ≤ 20 μl Whole blood: ≤ 13 μl | 20,000 results | 5 total: M-18D Diluent M-18R Rinse M-18CFL Lyse M18E E-Z cleanser M-18P Probe cleanser | Integrated thermal printer | 18
Mindray | BC 5000 | Tabletop | 23 3 histograms 4 scattergrams 12 research | WBC, RBC, HGB, MCV, PLT | 5-part | 4 USB ports, WiFi, LAN port | 40 | 15 μl for whole blood, capillary blood, prediluted | 20,000 results | 4 total: M-52D Diluent M-52-DIFF Lyse M-52LH Lyse Probe cleanser | external printer | 18
Mindray | BC 5150 | Tabletop | 25 3 histograms 4 scattergrams 24 research | WBC, RBC, HGB, MCV, PLT | 5-part | 4 USB ports, WiFi, LAN port | 60 | 15 μl for whole blood, capillary blood, prediluted | 250,000 results | 4 total: M-52D Diluent M-52-DIFF Lyse M-52LH Lyse Probe cleanser | external printer | 18

## Bedside Phys Mon
|  | Intended Use | Measured Parameters | Alarms | Display size (diagonal) | Operation Mode | Battery life (hours) | Connectivity | Warranty (months)
GE Healthcare | CARESCAPE B105 | Adult Pediatric Neonate | ECG (3, 5-lead), Resp, SpO2, NIBP, 2x temp, IBP Optional - C.O., EtCO2 | Audio & Visual Full capabilities | 25.6 cm 6 waveforms | Continuous | Rechargeable - 3 | RS-232, USB optional WiFi | 12
GE Healthcare | CARESCAPE B125 | Adult Pediatric Neonate | ECG (3, 5-lead), Resp, SpO2, NIBP, 2x temp, Optional - EtCO2, IBP | Audio & Visual Full capabilities | 30.7 cm 6 waveforms | Continuous | Rechargeable - 11 | USB, RS-232 | 12 24 - Moldova
Mindray | ePM 10/12 | Adult Pediatric Neonate | C.O., ECG (3, 5, 12-lead) ETCO2, IBP, NIBP, PR, RR, SpO2, body temp | Audio & Visual Full capabilities | 25.6 cm/30.7 cm | Continuous | Rechargeable - 2-4 | USB, RS-232 | 18
Mindray | iMEC 10/12 | Adult Pediatric Neonate | C.O., ECG (3, 5-lead) ETCO2, IBP, NIBP, PR, RR, SpO2, body temp | Audio & Visual Full capabilities | 26.4 cm/ 31 cm | Continuous | Rechargeable - 2-4 | USB, RS-232 | 18
Mindray | uMEC10 | Adult Pediatric Neonate | CO, ECG (3, 5, 12-lead) ETCO2, IBP, NIBP, PR, RR, SpO2, body temp | Audio & Visual Full capabilities | 26.4 cm | Continuous | Rechargeable - 4-8 | USB, RS-232 | 18
Philips | Efficia CM100 | Adult Pediatric Neonate | ECG (3- 5-lead), Resp, SpO2, NIBP, 2x temp, EtCO2, 2x IBP | Audio & Visual Full capabilities | 25.4 cm 8 waveforms | Continuous | Rechargeable - 9 | 2x USB wired LAN, ECG analog output, VGA | 12
Philips | Goldway G30E | Adult Pediatric Neonate | ECG (3- 5-lead), Resp, SpO2, NIBP, 2x temp 2x IBP, EtCO2 | Audio & Visual Full capabilities | 26.4 cm 8 waveforms | Continuous | Rechargeable - 2.5 | USB, VGA output HL7 EMR/HIS integration | 12

## Defibrillator
|  | Intended Use | Functions | Monitored Parameters | Max Energy Settings (Joules) | Shock Delivery Options | Alarm Capabilities | Battery Runtime (hours) | Warranty (months)
Mindray | BeneHeart D3 | Adult Pediatric | Manual defibrillation AED, non-invasive pacing synchronized cardioversion, monitoring | 3/5 lead ECG, respiration rate, NIBP, SpO2, ETCO2 | 360 J | Paddles Adhesive Pads | Audio & Visual Full capability | Rechargeable - 6 | 18
Mindray | BeneHeart D6 | Adult Pediatric | Manual defibrillation AED, non-invasive pacing synchronized cardioversion, monitoring | 12-lead ECG, respiration rate, NIBP, SpO2, ETCO2, Body Temp, IBP | 360 J | Paddles Adhesive Pads | Audio & Visual Full capability | Rechargeable - 6 | 18
Philips | Efficia DFM100 | Adult Pediatric | Manual defibrillation AED, non-invasive pacing, synchronized cardioversion, monitoring | 3-lead ECG Optional: NIBP, SpO2, ETCO2 (and respiration rate) | 200 J | Paddles Adhesive Pads | Audio & Visual Full capability | Rechargeable 2.5 hours monitoring 2 hours pacing 100 full energy shocks | 12

## Diagnostic ECG
|  | Intended Use | Number of Recorder Channels | Maximum Number of Leads | Intrepretative Analysis Program | Measurements | Display Screen (diagonal) | User Interface | Arrhythmia Detection | Defibrillator overload protection | Internal Storage | Digital Interface | Battery Runtime | Data Export | Patient ID Method | Integrated Printer | Warranty (months)
GE Healthcare | Mac 600* | Adult Pediatric | 4 | 12 | Marquette 12SL | Ventilation rate, P duration, PR interval, QRS duration, QT/c interval, P-R-T axis | Color 11 cm | Panel buttons | ✔ | ✔ | Unspecified, ECGs are saved on SD Card and transferred to a PC for detailed waveform viewing | 2 GB SD Card | 250 ECGs 3 hours w/o printing | PDF and XML | Manual | Yes | 12
GE Healthcare | Mac 2000 | Adult Pediatric Neonate | 12 | 12 | Marquette 12SL | Ventilation rate, P duration, PR interval, QRS duration, QT/c interval, P/QRS/T axis, RV5/SV1, RV5+SV1 | Color LCD 18 cm | 7 inch color display Full Keyboard | ✔ | ✔ | 200 ECGs | LAN, WiFi, 2x serial, modem, 2x USB, SD card | 100 ECGs not specified with printing 3 hours w/o printing | PDF and XML | Manual Barcode Scanner | Yes | 12 24 - Moldova
Hillrom | ELI 150c | Neonate Adult Pediatric | 12 | 12 | Welch Allyn VERITAS | HR, PR interval, QRS duration, QT/QTc interval, P/QRS/T axis, RV5/SV1, RV5+SV1 | Color LCD 1/4 VGA 320 x 240 pixels | Full keyboard | ✔ | ✔ | 40 ECGs (optional expansion to 200 ECGs, not included) | LAN, USB, RJ11 Optional (not included): WiFi, WLAN, GPRS, USBD | 150 ECG continuous printing 450 ECG w/o printing | XML Via ELI Link (optional, not included): PDF, DICOM | Manual Optional Barcode Scanner (not included) | Yes | 24
Mindray | Beneheart R3 | Adult Pediatric Neonate | 4 | 12 | U of Glasgow | HR, PR interval, QRS duration, QT/QTc interval, P/QRS/T axis, RV5/SV1, RV5+SV1 | Color LCD 13 cm | Panel buttons | ✔ | ✔ | 800 ECGs | LAN, WiFi, 2x USB, RJ45 | 500 ECGs 2 hours continuous recording 6 hours w/o printing | PDF and XML | Manual Optional Barcode Scanner | Yes | 18
Mindray | Beneheart R12 | Adult Pediatric Neonate | 12 | 12 | U of Glasgow | HR, PR interval, QRS duration, QT/QTc interval, P/QRS/T axis, RV5/SV1, RV5+SV1 | Color LCD 20 cm | Full keyboard | ✔ | ✔ | 800 ECGs | LAN, WiFi, 2x USB, RJ45 | 400 ECGs 1 hour continous recording 3.5 hours w/o printing | PDF and XML | Manual Optional Barcode Scanner | Yes | 18
Philips | PageWriter TC10** (stand-alone) | Adult Pediatric | 12 | 12 | Philips DXL ECG Algorithm; STEMI (ST segment elevation myocardial infarction) diagnostic aid for Culprit Artery | HR, RR interval, PR interval, QRS duration, QT/QTc interval, P, QRS, and T-wave axes | Color LCD 18 cm | Panel buttons | ✔ | ✔ | N/A | LAN | 300 ECGs 2 hours continuous recording 10 hours w/o printing | PDF and XML | Manual Optional Barcode Scanner (not included) | Yes | 36
Philips | PageWriter TC10** (Wireless w/Disclosure & Events, and Archive features) | Adult Pediatric | 12 | 12 | Philips DXL ECG Algorithm; STEMI (ST segment elevation myocardial infarction) diagnostic aid for Culprit Artery | HR, RR interval, PR interval, QRS duration, QT/QTc interval, P, QRS, and T-wave axes | Color LCD 18 cm | Panel buttons | ✔ | ✔ | 200 ECGs | LAN, WiFi, USB | 300 ECGs 2 hours continuous recording (6 lead) 10 hours w/o printing | PDF and XML | Manual Optional Barcode Scanner (not included) | Yes | 36
* Mac 600 specifically designed for traveling to patients in rural areas for ECG recording
** Used for diagnosing atrial fibrillation

## Laryngoscope
|  | Intended Use | Blade Style | Blade Sizes | Handle and Battery Type | Heavy-wall plastic or metal case | Handle Light | Warranty (months) | ISO 7376:2009 Compliance
Hill-Rom | Fiber Optic Macintosh Set (P/N 69696) | Adult Pediatric | MacIntosh | 1, 2, 3, 4 | "C" cell battery handle "AA" cell battery handle | ✔ | 2.5 V Halogen | 12 | ✔

## O2 concentrator
|  | Flow Range LPM | Oxygen Concentration | Sound level (dBA) | Alarms | Outlet Pressure | Battery | Warranty (months)
Philips | Everflo International OPI 230V EU | 0.5 to 5 | 93% ± 3% at 5 LPM | 45 | Oxygen percentage indicator Low: < 82%, Very Low: < 70% Power loss Low/High pressure Occlusion System malfunction | 5.5 PSI 38 kPA | ✖ | 12

## Mobile XRay
|  | Digital | Analog | Generator Power | DICOM | Display | Motor assisted drive | Wireless X-ray detector | Articulating arm with two pivot points | Battery life (full charge, exposures) | Exposure on main power | Cassette compartment | Warranty (months)
GE Healthcare | Brivo XR 115 |  | ✔ | 4 kW, 100 mA* | N/A | N/A | ✖ | N/A | ✔ | ✖ | Not specified | 6 | 12
GE Healthcare | Brivo XR 115 with Definium XR120 | ✔ | ✔ | 4 kW, 100 mA* | ✔ | N/A | ✖ | ✔ | ✔ | ✖ | Not specified | 6 | 12
GE Healthcare | Traveller XR120 with Definium XR120 | ✔ | ✔ | 32 kW, 400 mA | ✔ | 23.8" | ✖ | ✔ | Rotating, Sliding Column | 200 | Not specified | 5 | 12
GE Healthcare | Optima XR 220 AMX | ✔ |  | 30 kw, 300 mA | ✔ | 15" | ✔ | ✔ | Rotating, Sliding Column | 150 | ✔ | 3 | 12
GE Healthcare | Optima XR240 AMX | ✔ |  | 30kW, 300 mA | ✔ |  |  | ✔ | Rotating, Sliding Column | 100 | ✔ | 6 | 12
Mindray | Mobieye 700 | ✔ |  | 30 kW, 500 mA | ✔ | 19" | ✔ | ✔ | ✔ | 700 | ✔ | N/A | 18
Philips | MobileDiagnost wDR R2 | ✔ |  | 40 kW, 500 mA | ✔ | 17" | ✔ | ✔ | Rotating, Sliding Column | 500 | ✔ | Yes | 12
Siemens | Mobilett Elara Max | ✔ |  | 35 kW, 450 mA | ✔ | 17" | ✔ | ✔ | ✔ | 200 | ✔ | N/A | 12
Siemens | PolyMobil Plus |  | ✔ | 16 kW, 250 mA | N/A | N/A | ✖ | N/A | ✔ | N/A | ✔ | 8 | 12
* Brivo generator power is lower than ECRI recommendations of 15kW. This system should be adequate for routine chest and appendage x-rays. But may be suspectible to noise and blurring due to lower exposure times. It will have limited abdomen imaging capabilities.

## Otoscope
|  | Intended Use | Retinoscope Head | Ophthalmoscope Head | Proposed Disposable Specula | Throat Illuminator | Light Bulb | Battery Type | Warranty (months)
Hill-Rom | Halogen Otoscope (PN 22820) | Adult Pediatric | ✔ (optional) | ✔ (optional) | 4.25 mm - adult 2.75 mm - pediatric | ✔ | 2.5 V Halogen | 2x AA | 12

## Pulse Ox
|  | Measured Parameters |  |  |  |  | Measurement Limits and Accuracy |  |  |  |  | Standards Compliance
|  | peripheral capillary O2 saturation (SpO2) | Pulse Rate (PR) | Perfusion Index (PI) | Pleth Variability Index (Pvi) | Operation Mode | SpO2 detection range and accuracy | Pulse Rate range and accuracy | Physiologic Alarms | Communication and Data Export | Battery life | IEC60601 | ISO 80601-2-61 (Pulse Oximetry) | CE Mark | Warranty (months)
Mindray | PM-60 | ✔ | ✔ | ✖ | ✖ | Spot Check Continuous | Range: 0-100% No motion: 2% Motion: unknown Low perfusion: unknown | 18-300 BPM no motion: 3 BPM motion: 5 BPM Low Perfusion: unknown | Audio and Visual SpO2 - Low, High PR - Low, High | Data cable to PC | 24 hours* -Li-ion 36 hours* - AAA | ✔ | unknown | ✔ | 18
Masimo | RAD-G | ✔ | ✔ | ✔ | ✖ | Spot Check Continous | Range: 70-100% No motion: 2% Motion: 3% Low Perfusion: 2% | 25-240 BPM no motion: 3 BPM motion: 5 BPM Low Perfusion: 3 BPM | SpO2 - Low, High PR - Low, High RR - Low, High | ✖ | 24 hours* - Li-ion | ✔ | ✔ | ✔ | 24
Masimo | RAD-97 | ✔ | ✔ | ✔ | ✔ | Continuous | Range: 70-100% No motion: 1.5% Motion: 1.5% Low Perfusion: 2% | 25-240 BPM no motion: 3 BPM motion: 5 BPM Low Perfusion: 3 BPM | SpO2 - Low, High PR - Low, High RR - Low, High | WiFi, Bluetooth, Ethernet, USB Port | 4 hours* - Li-ion | ✔ | ✔ | ✔ | 24
* The battery-powered operating time of pulse oximeters is highly variable and fluctuates widely based on configuration and operation of the device. These estimates are based on the lowest display brightness and many features like WiFI turned off.

## Ultrasound
|  | Form Factor |  |  | Probe Types |  |  |  | Operating Modes |  |  | Clinical Applications
|  | Portable-laptop | Portable-tablet | Cart-based | Wireless | Flat linear Array | Convex Array | Phased/Sector Array | B-mode | Color Doppler | Power Doppler | Point of Care | Cardiac | General | Battery(> 30 min) | Warranty (months)
GE Healthcare | Versana Active | ✔* |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | Not specified | 12 36 - Moldova
GE Healthcare | Logiq E | ✔* |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | 12
GE Healthcare | Vivid Iq R3 | ✔* |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | 12
GE Healthcare | Logiq V2 | ✔ |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | 12
GE Healthcare | Logiq e Pro | ✔* |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | 12 36 - Moldova
Mindray | M6 | ✔ |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✖ | ✔ | ✔ | 18
Mindray | M7 | ✔ |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✖ | ✔ | ✔ | 18
Mindray | M9 | ✔ |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | 18
Philips | Lumify (3 probes) |  | ✔ |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✖ | ✔ | ✔ | ✔ | ✔ | 48 (system) 6 (probes)
Siemens | ACUSON Freestyle |  | ✔ |  | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✖ | ✔ | ✔ | 12
Siemens | ACUSON P500 | ✔ |  |  | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✖ | ✔ | 12
Siemens | ACUSON Juniper |  |  | ✔ | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | 12
Siemens | ACUSON NX2 |  |  | ✔ | ✖ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | ✔ | 12
*Cart available

## Ventilators
|  | Primary Available Modes | Indicated Patients | Tidal Volume (mL) | Max Inspiratory Flow (L/min) | Respiratory Rate (BPM) | CPAP/PEEP (cm H2O) | Max Pressure Support (cm H2O) | Inspiratory Time (seconds) | I:E Ratio | Inlet Gas Pressure Range (psi) | Integrated Air Compressor | Battery Runtime (minutes) | Alarm Capabilities | Warranty (months)
GE Healthcare | CARESCAPE R860 | A/C VC, A/C PC SIMV VC & PC CPAP/PSV, SBT | Adults Pediatrics Neonates | 2 to 2000 | 160 | 1 to 150 | 2 to 20 | 60 | 0.1 to 10 | 1:9 to 4:1 | 35 to 94 | Optional | 85 | full capabilities | 12 24 - Moldova
GE Healthcare | Falco 202 Evo | APCV (BILEVEL ST), APCV-TV, PSV (BILEVEL S), PSV-TV (Autoweaning), VC/VAC, VC/VAC BABY, V SIMV+PS, P SIMV+PS, SPONT, CPAP, APRV, NIV | Adults Pediatrics Neonates (optional) | 20 to 3000 | 160 | 4 to 150 | 3 to 50 | 80 | 0.04 to 9.6 | 1:10 to 4:1 | 40-86 | Optional | 240 | full capabilities | 12
Hill-Rom | Life2000** | Volume controlled ventilator: Control, Assist/Control, Assist | Adults | 0-2000 | 28 | 1 to 120 | 10 | 40 | 0.15 to 3.0 | N/A | 41 to 87 | Yes | 300-360 | Numerous adjustable alarm settings appropriate for this ventilator's capabilities | 12
Mindray | SV300 | A/C VC, A/C PC SIMV VC & PC CPAP/PSV, SBT | Adults Pediatrics | 20 to 2000 | 100 | 1 to 100 | 1 to 45 | 80 | 0.2 to 10 | 1:10 to 4:1 | 40 to 87 | Yes | 90 | full capabilities | 18
Mindray | Synovent E3 | A/C VC, A/C PC SIMV VC & PC CPAP/PSV | Adults Pediatrics | 40 to 2000 | 180 | 1 to 120 | 1 to 45 | 100 | 0.2 to 10 | 1:10 to 4:1 | 41 to 94 | Yes | 90 | full capabilities | 18
Philips* | V60+* | S (spontaneous), S/T (spontaneous/ timed), PCV (pressure control ventilation), AVAPS (average volume-assured pressure support), and CPAP (continuous positive airway pressure) | Adults Pediatrics (>20kg) | 0-3000 | 150 | 0 to 90 | 4 to 25 | 40 | 0.3 to 3.0 | n/a | 40 to 87 | Yes | 360 | full capabilities | 12
Philips* | E30* | S (spontaneous), S/T (spontaneous/ timed), PCV (pressure control ventilation), and CPAP (continuous positive airway pressure) | Adults Pediatrics (7 yo & >18kg) | n/a | 30 patient circuit 60 device inlet | 0 to 40 | 20 max | n/a | 05. to 3.0 | n/a | low flow only | n/a | No | Disconnect, Apnea, Low Minute Volume, Vent Inop, Pressure, Power failure | 12 device replacement only
* Both Philips ventilators are closer to low-mid-acuity ventilators and do not meet specs for critical care ventilators. This distinction is explained in detail within the contract launch summary and its appendix
Note: The Philips Respironics E30 ventilator does not have CE or FDA approval, and is provided globally for use under local emergency use authorizations, such as the FDA Emergency Use Authorization for ventilators, Health Canada Interim Order for use in relation to COVID-19, and waiver of CE marking, which authorize its use for the duration of the COVID-19 public health emergency, unless terminated or revoked (after which the products may no longer be used).
** Hill-Rom Life 2000
Brief summary of the Life2000 design and functionality
This is a volume controlled predominantly noninvasive ventilator (90% of use) capable of some invasive ventilation if following a low PEEP/high FiO2 therapy
It comes with a priority universal connector for nasal cannula, mask, ETT
Ventilator is most appropriate for treating patients before the need of an ICU ventilator and then post-ICU. This would not replace an ICU ventilator but it may help with ventilator fleet management especially with ICU ventilator in high demand.
However, clinical judgement about when to transition a patient from the Life2000 to an ICU vent may be a challenge. Indicators are unstable SpO2 and deterioration of RR
PEEP up to 10 cm, PIP up to 40 cm – low PEEP is a limiting factor for invasive ventilation, it can deliver up to 90% O2
6 hour runtime on full battery charge, can run while battery is charging, not hot swappable
Ventilator is pneumatically driven – dependent on a pressurized O2/air source (wall air or compressed gas) because it doesn’t have an internal blower
The ventilator can come with an air compressor option – I asked Hill-Rom to propose this as an optional feature
Total delivered volume is different than the amount set by the user via the ventilator settings: – the total amount equals the volume set by the user (O2) and about two times that set volume of air entering the patient circuit as air entrainment - this requires TRAINING to user to understand this difference
Has a number of alarms high/low PIP, high/low RR, high circuit pressure, low gas inlet pressure etc. but no as comprehensive as a ICU vent
Note on B5: Schabowsky, Christopher: This is not an ICU vent with notable drawbacks. It is primarily a non-invasive ventilator. See description below.

## Vital Signs
Measured Parameters
|  | Intended Use | NIBP | Body Temp | Pulse Rate | SpO2 | Operation Mode | Physiologic Alarms | Battery life (hours) | Connectivity | Warranty (months)
GE Healthcare | CARESCAPE V100 | Adult Pediatric Neonate | ✔ | ✔ | ✔ | ✔ | Spot Check Continuous | Systolic High/Low Diastolic High/Low Pulse rate High/Low SpO2 High/Low | 11 | USB, RS-232 | 12 24 - Moldova
Hillrom | ProBP 2000 | Adult Pediatric | ✔ | ✖ | ✔ | ✖ | Spot Check | N/A Spot Check Only | Requires four AA batteries; life not specified | ✖ | No warranty
Mindray | VS600 | Adult Pediatric Neonate | ✔ | ✔ | ✔ | ✔ | Spot Check | N/A Spot Check Only | 11 low capacity batt. 22 high capacity batt. | ✖ | 18
Note on B4: Schabowsky, Christopher: This is only a blood pressure cuff. Hill-Rom may propose more appropriate models soon.

## Device Type Tracking
| Analyzers Hematology | Analyzers POC Blood Gas | Defibs | ECGs | Food Transport Trolley | Laryngoscopes | Mobile X-Rays | Needle Cutter | Otoscopes | Oxygen Concentrator | Oxygen Flow Meters | Patient Beds | Phy Mons | Plasma Thawer | Pulse Ox | Suction Devices | Steam Sterilizers | Thermometer Digital, IR | Tool table mobile | Ultrasound | Ventilators | Vital Signs Monitors | Wheelchairs
ECRI Tech Specs | ✔ | ✔ | ✔ | ✔ |  | ✔ | ✔ |  | ✔ | ✔ |  |  | ✔ |  | ✔ |  |  |  |  | ✔ | ✔ | ✔
Proposals Received
GE |  |  | ✔ | ✔ |  |  | ✔ | [5 empty] | ✔ | [6 empty] | ✔ | ✔
Hill Rom |  |  |  | ✔ |  | ✔ |  |  | ✔ |  |  | ✔ | [8 empty] | ✔* | ✔
Masimo | [14 empty] | ✔ | [5 empty] | ✔*
Mindray | ✔ |  | ✔ | ✔ |  |  | ✔ | [5 empty] | ✔ |  | ✔ |  |  |  |  | ✔ | ✔ | ✔
Philips |  |  | ✔ | ✔ |  |  | ✔ |  |  | ✔ |  |  | ✔ | [6 empty] | ✔ | ✔*
Siemens |  | ✔ |  |  |  |  | ✔ | [12 empty] | ✔
Proposals Incoming
Laerdal | [15 empty] | ✔
Linet | [11 empty] | ✔
Potential Proposals
Stryker | [11 empty] | ✔
Table Key
High demand (> 5 countries) device category met with multiple vendor bids
High demand (> 5 countries) device category needs more vendor bids
High demand, no vendor bids
| * | Not an ICU ventilator