## Emergency Car and components
| Items | Quantity | Unit Price | Total Price | Delivery/ Lead Time | Country of Origin | Notes
1 | ambulance car c type (intensive care mobile unit) | 30 | EUR 69,500 (DAP Batumi) | 2085000 | 35 days from signed PO | Kosovo | Opel Movano L3H2 with specifications as per UNOPS Pristina tender (specifications available)
|  |  | EUR 64,900 (DAP Batumi or Poti) | 1947000 | 35 days from signed PO | Kosovo | Renault Master model year 2020 L3H2 with specifications as per UNOPS Pristina tender (specifications available)
|  |  | $48,500 (EXW) | 1455000 | 90-150 days | Turkey | Quote from Mobilambulans, Turkey; Ford Transit 350L High Roof 170 HP Diesel 2020 Model Manual Transmission meeting requirements in "Specs Ambulance" Tab (specifications available)
|  |  | $53,500 (EXW) | 1605000 | 90-150 days | Turkey | Quote from Mobilambulans, Turkey; Volkswagen Crafter 2.0 TDI Mid Chassis High Roof 140 HP Diesel 3.5T 2020 Model Manual Transmission meeting requirements in "Specs Ambulance" Tab (specifications available)
|  |  | $55,010 (FCA) | 1650300 | 45 days production time | USA | UN Web Buy Plus, Ford Transit 150 Advanced Ambulance, 2WD, 3.7L Petrol Auto, Medium Roof, LHD (specifications available, and different medical equipment options available)
|  |  | 31000 | 930000 | 150 Day Delivery to Tbilisi | Italy | Fiat Ducato Ambulance - Price does not include full medical equipment, final quotation will be provided early next week.
2 | Motorcycle ambulances | 50 | $3,945 (50+ units, FCA) | 197250 | 10 weeks | Netherlands | LTA available, consideration should be given to the feasibilty and quality, and matched with Georgian requirements
|  |  | 5900 | 295000 |  | South Africa | Company: eRanger
3 | Motocycle for ambulance firs responders with helmets for ambulance firs responders | 50 |  |  |  |  | Additional information/specifications required - whether specific helmet type or different motorcycle than above?
4 | CPR machine lucas 2 | 150 | 16000 | 2400000 | 6 Weeks | USA | LUCAS 3 system also available with marginal price increase
5 | Rec. camera for ambulance patient cabine | 400 | 170 | 68000 | 6 weeks | USA | Wide angle 170°. Can also be used as dashcam
6 | Tablet for Ambulance Emergency Medical Care Car | 400 | 300 | 120000 | 6 Weeks | China | Lonovo yoga Tab, with extended battery,
7 | Body worn camera for ambulance group | 700 | 200 | 140000 | 6 Weeks | China

## Specs
Diagnostics | Min Specs
nCov-2019 Rapid detection kit (IgM/IgG) | Utilizes human whole blood (preferably peripheral), serum, or plasma· Built-in control (sensitivity) ≥ 90%, specificity ≥ 95%, number of samples used for clinical evaluation ≥ 100; · Absence of cross-reactions: human coronavirus panel, HBV, HCV HIV-1 HIV-2 Adenovirus, Parainfluenza virus 1-4, Influenza A, Influenza B, Respiratory syncytial virus, Rhinovirus· Preferably: Follow up with additional positive and negative control samples, sample collection kit and cross-examination: Human Metapneumovirus (hMPV) Chlamydia pneumoniae Streptococcus pneumoniae Mycobacterium tuberculosis Mycoplasma pneumoniae EB Virus, Enterovirus 71· Used in rapid, qualitative and differential detection of IgG and IgM antibodies· Delivers clinical results between 10 and 20 minutes at the point-of-care· Visual interpretation of results· No special equipment needed· CE-Marked, IVD, Safe Packaging, Labeling Method: Immunochromatographic· STANDARD: to be approved by FDA
nCov-2019 Rapid detection kit (Antigen)
nCOV-19 Real-time PCR (with extraction kits) | The kit should be multiplex assay with capability of detection of at least two target genes and internal extraction control. Negative and positive PCR controls as well as master mix should be included in the kit.The limit of detection with 95% confidence interval (LoD 95) for screening / confirmatory assays should be at least 8 / 10 copies per reaction.The specificity of the kit should be tested to demonstrate the analytical specificity to Coronavirus SARS-CoV-2. It should not show the cross amplification with other human pathogens including human coronaviruses (229 E, HKU1, MERS, SARS, NL63, OC43), Influenza, RSV and other respiratory viruses.The limit of detection with 95% confidence interval (LoD 95) for screening / confirmatory assays should be at least 8 / 10 copies per reaction.The specificity of the kit should be tested to demonstrate the analytical specificity to Coronavirus SARS-CoV-2. It should not show the cross amplification with other human pathogens including human coronaviruses (229 E, HKU1, MERS, SARS, NL63, OC43), Influenza, RSV and other respiratory viruses.
ECG 12 lead | 12-lead ECG with colour screen, with ECG recording manual and automatic function, with battery, charging time no more than 4 hours. With the ability to print on paper not less than 80 mm, with ECG cable, with multiple electrodes. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan)
patient monitor for vital signs | Patient monitor, with colour display, for adults, pediatric and neonatal patients options, with wall or shelf mounting device, (ECG, NIBP, SpO2, T, HR, P), ECG chanels: I, II, III, aVR, aVL, aVF, V, Equipped with a set of electrodes, transportation battery( 12W და 220W Charging device), storage working duration at least 5 hours. ISO, CE Certificate existence, with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan)
portable oxygen concentrator | Concentrator with a capacity not less than 5 liters of oxygen per minute, with flow regulation, with inhalation function, oxygen production not less than 93%
Pulse oximeters | Portable Pulse oximeter, with color screen, on display - pulse, percentage of oxygen, with battery, with hanger and protective bag
Infrared thermometers | Non-contact thermometer. Measurement of body temperature. Hygienic and safe-measure without skin contact. Measurement in <1 second Infrared measurement technology Display in Celsius or Fahrenheit Free high temperature alarm setting with battery
portable transport ventilator for adults pediatric and neonatal patients, aproved for use helicopters and airplanes | transport ventilator for adults, pediatric and neonatal patients (IPPV, SIMV, CPAP, PEEP mode, pressure controllated ventilation mode) peep from 0 to no less than 20 mBar (cmH20), with battery (DC12v and AC220v charging device). Multiple flow sensor, the basic and a spare countour for adult and neonatal patients, (countour with test lung and face mask) brand new, color display, wall mounting, with transportation bag, with central and local pipe of the oxygen, with own oxygen cylinder, working temperature range at least from ‐18 c to +50 c‐ , the oxygen concentration mixed and 100% , storage battery at least 5 hours working time, charging time not more than 5 hours, must have a certificate of ISO 10651-3, CE certificate, agreeng document to EN 1789. ventilator should have certificate for using at air transport.
Defibrilator with patient monitoring system for adults and pediatric peds | External biphasic Defibrillator, output energy levels from 5 to not less than 270 J with adult and pediatric integrated paddles, with battery, with colour display, ECG, SpO2, P index, external pacing , battery operation time in full charge at least 300 min. ISO, CE Certificate existence, with wall mounting device. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan)
ECG 12 lead | 12-lead ECG with colour screen, with ECG recording manual and automatic function, with battery, charging time no more than 4 hours. With the ability to print on paper not less than 80 mm, with ECG cable, with multiple electrodes. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan)
portable oxygen generation system (pogs) | Equipment with the ability to quickly transport and arrange oxygen in case of emergency, in order to be used in field missions. It should be possible to supply oxygen to the patient directly from the device, as well as to pump oxygen cylinders. It should be possible to provide oxygen for the breathing ventilator with this device.
patient monitor for vital signs | Patient monitor, with colour display, for adults, pediatric and neonatal patients options, with wall or shelf mounting device, (ECG, NIBP, SpO2, T, HR, P), ECG chanels: I, II, III, aVR, aVL, aVF, V, Equipped with a set of electrodes, transportation battery( 12W და 220W Charging device), storage working duration at least 5 hours. ISO, CE Certificate existence, with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan)
portable oxygen concentrator | Concentrator with a capacity not less than 5 liters of oxygen per minute, with flow regulation, with inhalation function, oxygen production not less than 93%
Pulse oximeters | Portable Pulse oximeter, with color screen, on display - pulse, percentage of oxygen, with battery, with hanger and protective bag
Field Hospital tent sets for bioisolation patients | The tent should be at least 30 sq.m. with a floor, lighting system, heating and cooling system, if necessary with a compressor, the possibility of integration with another tent, in order to manage infectious patients, in case of need the patient should be isolated in the tent.
isolation patient transportation capsule | Transport isolation capsules, especially for transportation of patients with dangerous infections, should be available medical manipulations with the patient placed in the capsule, if necessary, without unisolation, with a set of appropriate filters and a gas circulation device.
CPR machine lucas 2 | Automatic cardiac resuscitation device, the ability to automatically perform chest compressions without the involvement of medical personel, frequency not less than 100 pressures per minute
Bone Catheter for adults and pediatric patients | Bone catheter for adults and pediatric patients (1000 pieces for adults and 500 pieces for pediatric)
Rec. camera for ambulance patient cabine | Recorder with data storage and transmission (wifi / lan). It should be possible to start / finish recording manually as well as set it to automatic mode
Tablet for Ambulance Emergency Medical Care Car | Tablets for ambulances, with the appropriate standard of drop, vibration safety, GPS, WIFI function, with the ability to support at least 4 G Internet, equipped with charger, car mount device
Body worn camera for ambulance group | body worn recording camera with data storage and transmission capability,with manually start / finish recording function, with the ability to store data for at least 24 hours

## Specs Ambulance
N | Name | ambulance car
1 | Minibus-type ambulances | New -Not used (Mileage -No more than 1000 km), manufacturing yaer - No earlier than 2019
2 | Vehicle body | Full body
3 | Vehicle Color | white
4 | wheel location | left
5 | Four Wheel Drive or/and on/off of axle, locking differential, low transmit | set
6 | Power Steering hydraulic | set
7 | ABS system | set
8 | A mandatory requirement | directional stability system (ESP, ESC or/and another one)
9 | A mandatory requirement | audio system Radio
10 | Protective cushion (airbag) in the driver's cab | 2 or more
11 | Engine nominal power (not less than euro 5 standarts) | not less than 125 hp
12 | Fuel Type | diesel
13 | fuel tank | not lees 70 Lt
14 | Gearbox | Mechanics 7 step( 6+1, including rear moving step)
15 | protective shield of engine lower part | set
16 | Rear salon length | not less than 3.45 meters
17 | Rear salon width | Not less than 1.7 meters, width between the rear wheel arches at least 1.35 meters
18 | The height of the rear cabin (from floor) | not lees 1,85 meters
19 | Side Door width of the patient's booth (From the opposite side of the steering wheel), with glass (opening / closing function) | not lees 1,1 meters
20 | Rear door Height | not lees 1,5 meters
21 | The width of the rear doorway | not lees 1,3 meters
22 | Rear door type | 2 (two) aside open door
23 | number of doors | 5
24 | ground clearence | not lees 0.15 meters
25 | not less R 16 diameter tires. ( including same size spare tire) * | set
26 | Patient cabin’s each doors should be provided with support device, to prevent the arbitrary closing of the door. | set
27 | Storage battery | 2 (two) pieces or more, each not less 90 Ampère
28 | Storage battery charging spare device | set
29 | Design
30 | Thermal insulation | set
31 | Front and Rear salons heating / cooling system. For regulator In rear saloon electric sensor board | set
32 | Front and Rear salons dividing partition with sliding window (window should have a lock with fixator, Front and rear salon must have a full body construction) sizes with agreeng buyer | set
33 | Ceiling and wall coating (Heat, sound and hydro insulation) | set
34 | Extinguisher (Car) for Front and Rear salons | set
35 | Floor (antibacterial) water resistant insulation, floor cover anti static and must overlap on the wall and shelves not less by 5 cm | set
36 | Rear doors with semi-transparent covers glass. Rear door should have installed steps (level), so that Med- Staff did not find it difficult to access from the back door. | set
37 | Rear salon Glass with semi-transparent covers | set
38 | Handle stick on entire length of rear cabin ceiling (solid construction) | set
39 | Infusion system hanger, (Built-in ceiling or hanging-on) | not less 4 set
40 | Number of seats in front cabin (including the driver) | 3
41 | interior of rear saloon should be made with "fiber glass" or "ABS plastik" material, with hydroisolation, made with fire-proof, desifenctant liquid durable material | set
42 | Oxygen system, cylinders with mounting accessories, gas regulator removable crunch, with regulators | set, total volume of oxygen not less than 40 liters, number of cylinders agreeng with buyer
43 | Oxygen volume regulator | 1
44 | Oxygen pressure regulators | not less 2
45 | Cabinets, shelves, store boxes, hooks ** | set
46 | Chairs and Seats in the rear cabin for med-personnel: seat with folding chair, in parallel to the stretcher, with safety belts | 1
47 | med-personnel's rotating seat with handles (turn around 90 degrees both sides), with safety belt, placed parallel to the stretcher, at top part of the stretcher. | 1
48 | Folding chair seat (Seat backrest on the Salons partition, Not less than 0.4 m wide) with safety belt. | 1
49 | bin for medical waste, capacity non less 5 liters | 1
50 | Electrical equipment, power
51 | One unit fleshing signal on the roof front side with diod fleshing, with outside transparent glasses (with side illumination), fleshing signal with blue and white light (with function projector), with animation remote control to change lighting options. working voltage -12 wolt, protection class not less than IP 65, length not less than 120 cm, with ce, R65 and SAEJ845 certificates, Two units on front bumper with blue-white LED-Diod technology, 3 units on both sides and 2 units at the back. with airtight performance, One unit fleshing signal on the roof rear side. | set
52 | Loudspeaker system at least 150 watts, Power output 115 dB, "loudspeaker" built-in flashlights or other place,, hermetic performance, at least nine standard regimens of voice, power supply constant electricity 12 volt, remote control integrated with fleshing signal. | set
53 | Ambulances vehicle relevant ceiling lighting with white color light (for possibility of good observation patient's vital signs: skin color, eye pupil reflex, etc.) | set
54 | Spot lights in celling (for stretcher) | 2 or more
55 | Patient cabin lighting and heating-cooling system Control Board should be installed on proper location, (not to make discomfort for the patient) | 1
56 | 12-220 voltage inverter (not less than 2 KWT) | set
57 | Back run headlight | 1
58 | front and rear fog lamps | set
59 | possibility of connection between driver and rear salon
60 | Back run voice signal | set
61 | 220 volt outlet in rear salon (european standard) | not less than 6
62 | 12 volt outlet in rear salon not less than 4 and in front salon not less than 2 | not less than 6
63 | Outside outlet for 220 voltage supply internal operation, with protective sheet and connecting cable 15 meters. | 1
64 | Ventilation set of the rear cabin | set
65 | car doors central locker, key with remote, with posibility to open rear and front cabines independently, alarm signal if door is open or not fully closed | set
66 | fuel tank hood with locking function | set
67 | video recorder, mounted in the drivers cabine, watching front area, vision not less than 150 degrees, memory not less than 60 gb, capacity at least 24 hours of video recording, with a minimum recording quality of 1920x1080, 30 fps (speed), equipped with GPS and Wi-Fi function | 1
68 | parking control system not less than 8 sensor and rear view camera | set
69 | Immobilizer Utilities
70 | Stretcher with receiver device, for adults . With its lower stand (made with rustproof material) with Height control mechanism, also lower stand with Height control mechanism, for stretcher download, to have appropriate movment capacity. lower stand - Possibility of width movement in the cabin. lower stand with damping system. stretcher with folding tires (locking device on wheels), with their guides, at least three belt for patient, Load capacity not less 220 kg; Mattress dimensions: no less than 55 cm X200 cm X85 cm (width, length, height), inaccuracy ± 5 cm; Weight: not more than 40 kg. made with desifenctant liquid durable material | set
71 | Stretcher infusion tripod (set on / set off function) | 1
72 | Spare Stretcher, (Folding, with an opportunity to get seat shape, with tires, wheel locks, with spetial equipment for moving on the stairs) comfortably fixed, not to create discomfort to the patient. | 1
73 | vacuum matress with appropriate suction | 1
74 | Immobilization shield (Patient safety belts type spyder, head fixator) The minimum length 1, 70 m, load capacity not less then 150 kg. | 1
75 | Medical Equipment and Supplies
76 | transport ventilator for adults, pediatric and neonatal patients (IPPV, SIMV, CPAP, PEEP mode, pressure controllated ventilation mode) peep from 0 to no less than 20 mBar (cmH20), with battery (DC12v and AC220v charging device). Multiple flow sensor, the basic and a spare countour for adult and neonatal patients, (countour with test lung and face mask) brand new, color display, wall mounting, with transportation bag, with central and local pipe of the oxygen, with own oxygen cylinder, working temperature range at least from ‐18 c to +50 c‐ , the oxygen concentration mixed and 100% , storage battery at least 5 hours working time, charging time not more than 5 hours, must have a certificate of ISO 10651-3, CE certificate, agreeng document to EN 1789. ventilator should have certificate for using at air transport. | 1
77 | Patient monitor, with colour display, for adults, pediatric and neonatal patients options, with wall or shelf mounting device, (ECG, NIBP, SpO2, T, HR, P), ECG chanels: I, II, III, aVR, aVL, aVF, V, Equipped with a set of electrodes, transportation battery( 12W და 220W Charging device), storage working duration at least 5 hours. ISO, CE Certificate existence, with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 1
78 | External biphasic Defibrillator, output energy levels from 5 to not less than 270 J with adult and pediatric integrated paddles, with battery, with colour display, ECG, SpO2, P index, external pacing , battery operation time in full charge at least 300 min. ISO, CE Certificate existence, with wall mounting device. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 1
79 | Infusion Pump for infusion solutions, with battery ISO, CE Certificate existence with wall mounting device | 1
80 | Pump, (2 syringes one of them not less 50 ml syringe capacity), with battery, ISO, CE Certificate existence, with wall mounting device | 1
81 | Aspirator with 1 liter jar or more, at least 3 mode of pumping, with battery, pump tube length at least 1.5 meters, with the negative pressure indicator. with wall mounting device, ISO, CE Certificate existence | 1
82 | Medical bag (Laryngoscope with set of straight and curved blades 0,1,2,3,4 sizes, ambu bag for adult, pediatric patients, oxygen cylinder 0.5 - 1 l with reducer, Medicines and medical goods storage) | 1
83 | Oxygen Device: oxygen distribution with flowmeter, Oxygen Moisture glass, Oxygen outlet 3-4 (DIN 25 standard) | set
84 | portable Oxigen cylinder with bag, with reducer and volume control set (volume not less 2.4 L) | 1
85 | kendrick fixator | 1
86 | portable refrigirator 10 liter capacity, temperature range not less -5 to +10, with posibility to fix in the patient salon | 1
87 | leg traction fixator | 1
** Spare tire and other equipment should be installed outside of the cabin, so that their use does not pose an inconvenience to the patient.
**All shelves and boxes should be fixed to each other, to reduce noise and ensure the uninterrupted use. All shelves and boxes should be screwed-on in the floor, to be able of their removing in case of the works will be required on the vehicle. Cabinets, shelves, storage should be open or closed, with drawers, with rendered or common doors, with fixator shutters, made with organic glass or solid material. Of sufficient size and configuration to provide all necessary equipment, medical consumables and medical devices placed in storage. All equipment must be readily accessible to personnel, even while loaded.Cabinets should be able closed firmly with lock (lock holder) to prevent inadvertent movement of the door. Cabinets, shelves, brackets and storage layout and dimensions should be pre-agreed with the buyer.
Emergency vehicles should be marked on with distinctive sign / symbol, which should be agreed in advance to the buyer.

## Delivery exp. (hidden)
Procurement Plan |  | Immediate Procurement: % of Estimated Needs by September 2020 | Second Procurement: % of estimated need by December 2020 | Medium term (1 and 2 Quarter of 2021)
0 | High Priority | 50% | 50%
0 | High Priority | 50% | 50%
0 | High Priority | 80% | 20%
0 | High Priority | 80% | 20%
0 | Medium Priority for equipment as it should be refurbished first | 20% | 30% | 50%
0 | High Priority | 100%
0 | High Priority | 80% | 20%
0 | High Priority | 60% |  | 40%

## Data Validation (hidden)
Selection:
Available through existing LTA/Contracts
Available stock
Not available