## ems georgia
| technical requirment | No of units requested - GEO EMS
Infrared thermometers | Non-contact thermometer. Measurement of body temperature. Hygienic and safe-measure without skin contact. Measurement in <1 second Infrared measurement technology Display in Celsius or Fahrenheit Free high temperature alarm setting with battery | 500
portable transport ventilator for adults pediatric and neonatal patients, aproved for use helicopters and airplanes | transport ventilator for adults, pediatric and neonatal patients (IPPV, SIMV, CPAP, PEEP mode, pressure controllated ventilation mode) peep from 0 to no less than 20 mBar (cmH20), with battery (DC12v and AC220v charging device). Multiple flow sensor, the basic and a spare countour for adult and neonatal patients, (countour with test lung and face mask) brand new, color display, wall mounting, with transportation bag, with central and local pipe of the oxygen, with own oxygen cylinder, working temperature range at least from ‐18 c to +50 c‐ , the oxygen concentration mixed and 100% , storage battery at least 5 hours working time, charging time not more than 5 hours, must have a certificate of ISO 10651-3, CE certificate, agreeng document to EN 1789. ventilator should have certificate for using at air transport. | 50
Defibrilator with patient monitoring system for adults and pediatric peds | External biphasic Defibrillator, output energy levels from 5 to not less than 270 J with adult and pediatric integrated paddles, with battery, with colour display, ECG, SpO2, P index, external pacing , battery operation time in full charge at least 300 min. ISO, CE Certificate existence, with wall mounting device. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 350
ECG 12 lead | 12-lead ECG with colour screen, with ECG recording manual and automatic function, with battery, charging time no more than 4 hours. With the ability to print on paper not less than 80 mm, with ECG cable, with multiple electrodes. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 350
ambulance car b type |  | 50
ambulance car c type (intensive care mobile unit) |  | 30
portable oxygen generation system (pogs) | Equipment with the ability to quickly transport and arrange oxygen in case of emergency, in order to be used in field missions. It should be possible to supply oxygen to the patient directly from the device, as well as to pump oxygen cylinders. It should be possible to provide oxygen for the breathing ventilator with this device. | 90
patient monitor for vital signs | Patient monitor, with colour display, for adults, pediatric and neonatal patients options, with wall or shelf mounting device, (ECG, NIBP, SpO2, T, HR, P), ECG chanels: I, II, III, aVR, aVL, aVF, V, Equipped with a set of electrodes, transportation battery( 12W და 220W Charging device), storage working duration at least 5 hours. ISO, CE Certificate existence, with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 20
portable oxygen concentrator | Concentrator with a capacity not less than 5 liters of oxygen per minute, with flow regulation, with inhalation function, oxygen production not less than 93% | 20
Pulse oximeters | Portable Pulse oximeter, with color screen, on display - pulse, percentage of oxygen, with battery, with hanger and protective bag | 700
Set of Tetra standard hand Radio station |  | 490
Set of Tetra standard Car Radio station |  | 380
Tetra system dispatch control and monitoring set |  | 43
Field Hospital tent sets for bioisolation patients | The tent should be at least 30 sq.m. with a floor, lighting system, heating and cooling system, if necessary with a compressor, the possibility of integration with another tent, in order to manage infectious patients, in case of need the patient should be isolated in the tent. | 4
isolation patient transportation capsule | Transport isolation capsules, especially for transportation of patients with dangerous infections, should be available medical manipulations with the patient placed in the capsule, if necessary, without unisolation, with a set of appropriate filters and a gas circulation device. | 15
Motocycle for ambulance firs responders with helmets for ambulance firs responders |  | 50
CPR machine lucas 2 | Automatic cardiac resuscitation device, the ability to automatically perform chest compressions without the involvement of medical personel, frequency not less than 100 pressures per minute | 150
Bone Catheter for adults and pediatric patients | Bone catheter for adults and pediatric patients (1000 pieces for adults and 500 pieces for pediatric) | 1500
Rec. camera for ambulance patient cabine | Recorder with data storage and transmission (wifi / lan). It should be possible to start / finish recording manually as well as set it to automatic mode | 400
storage for servers |  | 1
server |  | 4
Tablet for Ambulance Emergency Medical Care Car | Tablets for ambulances, with the appropriate standard of drop, vibration safety, GPS, WIFI function, with the ability to support at least 4 G Internet, equipped with charger, car mount device | 400
Body worn camera for ambulance group | body worn recording camera with data storage and transmission capability,with manually start / finish recording function, with the ability to store data for at least 24 hours | 100
TOTAL COST:

## consumable goods
No of units requested - GEO EMS
Gloves | 1000000
Gowns | 2000000
Goggles | 50000
Surgical masks | 3000000
N95 respirator masks | 300000
N95 mask fit test kit | 5
Antibacterial liquid soap (liters) | 50000
Alcohol-based hand rub (70%, liters) | 100000
Chlorine-based cleaning solution for surfaces | 10000

## Primary health care
| technical requirment | No of units requested - GEO EMS
ECG 12 lead | 12-lead ECG with colour screen, with ECG recording manual and automatic function, with battery, charging time no more than 4 hours. With the ability to print on paper not less than 80 mm, with ECG cable, with multiple electrodes. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 950
patient monitor for vital signs | Patient monitor, with colour display, for adults, pediatric and neonatal patients options, with wall or shelf mounting device, (ECG, NIBP, SpO2, T, HR, P), ECG chanels: I, II, III, aVR, aVL, aVF, V, Equipped with a set of electrodes, transportation battery( 12W და 220W Charging device), storage working duration at least 5 hours. ISO, CE Certificate existence, with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 950
portable oxygen concentrator | Concentrator with a capacity not less than 5 liters of oxygen per minute, with flow regulation, with inhalation function, oxygen production not less than 93% | 950
Pulse oximeters | Portable Pulse oximeter, with color screen, on display - pulse, percentage of oxygen, with battery, with hanger and protective bag | 1300
TOTAL COST: