| PURPOSE OF USE |  |  |
| --- | --- | --- |
| 1 | Clinical or other purpose | Ventilators designed to provide support to patients in Emergency Transport |
| 2 | Level of use (if relevant) | Health post, health center, Emergency Transport |
| 3 | Overview of functional requirements | Dispenses a controlled mixture of oxygen and air to the patient Gives artificial respiratory support as necessary Fully alarmed with all necessary monitors for continuous operation in transport environment Includes compressor and humidifier Reusable, sterilizable patient masks and / or connectors Suitable for all ages and body weights of patient |

| TECHNICAL CHARACTERISTICS |  |  |
| --- | --- | --- |
| 4 | Detailed requirements | The equipment should be configured to have pressure and flow triggers as default. Flexible breathing circuit a control system, monitors, and alarms. Microprocessor controlled Time, flow, volume and pressure adjustments with automatic compensation. Automatic self-test and sensor calibration. Automatic leakages compensation. Back-up ventilation in case of apnea during spontaneous mode of ventilation. Communication input and output ports to receive and transfer data, at least RS-232 in/out connections should be available. The equipment should be able to operate with either high and low pressure oxygen sources. The equipment should be able to operate with central gas supply and with medical gas cylinder sources. Type of gasses supported at least: Medical Oxygen and Medical Air. |
| 5 | Displayed parameters | LCD/ TFT color graphic high contrast display. Monitor parameters: • Oxygen concentration; • Tidal exhaled volume; • Pressure peaks and mean airway pressure; • Respiratory rate; • PEEP – From 0 to 20 mBar (cmH20) (at lest); • Exhaled minute volume; • Mean airway pressure; • Occlusion pressure detection; • Air and oxygen pressure; • Spontaneous ventilation; • Leak percentage; • Internal battery charge status. Real time flow and pressure curves monitoring and visualization, at least 3 simultaneously. |
| 6 | User adjustable settings | The following variables should be controllable by the operator: a) Tidal volume up to 1,000 mL. b) Pressure (inspiratory) up to 80 cm H20 c) Volume (inspiratory) up to 120 L/min d) Respiratory rate: up to 60 breaths per minute. e) SIMV Respiratory Rate: up to 40 breaths per minute. f) CPAP/PEEP up to 20 cm H2O. g) Pressure support up to 45 cm H2O. h) FiO2 between 21 to 100 % i) Inspiratory and expiratory times up to at least 2 sec and 8 sec respectively j) I:E Ratio at least from 1:1 to 1:3. 2 Modes of ventilation: a) Volume controlled. b) Pressure controlled. c) Pressure support. d) Synchronized intermittent mandatory ventilation (SIMV) with pressure support. e) Assist / control mode f) CPAP/PEEP Alarms required: FiO2, minute volume, pressure, PEEP, apnoea, occlusion, high respiration rate, disconnection System alarms required: power failure, gas disconnection, low battery, vent inoperative, self diagnostics If alarm silencing feature is incorporated, it must be temporary and clearly displayed when activated Air and externally supplied oxygen mixture ratios fully controllable Inlet gas supply (O2) pressure range at least 35 to 65 psi Medical air compressor integral to unit, with inlet filter |
| PHYSICAL/CHEMICAL CHARACTERISTICS |  |  |
| 7 | Components | External oxygen supply connection to be secure but easy to fit and release Case to be hard and splashproof Oxygen tube adapter must be compatible with central oxygen system and portable cylinder Supplied in protective case for clean storage and safe transport Panel settings protected from accidental operation Whole unit to be easily portable by hand Unit to be stable when tabletop mounted Controls and displays to be easily visible from the front of the unit in low light levels Mounting device in Emergency Transport Weight not higher than 7 Kg. |
| 8 | Mobility, portability | Equipment designed for easy and quick transportation. |
| UTILITY REQUIREMENTS |  |  |
| 10 | Electrical, water and/or gas supply (if relevant) | Power input to be ************* fitted with ********** compatible mains plug Maintenance-free rechargeable battery backup operation for minimum 5 hour in the event of mains power failure Low battery indicator required to alert user when on-board battery is low Battery to be recharged automatically when connected to mains power supply Voltage corrector / stabilizer / UPS to allow operation at ± 30% of local rated voltage and one hour operation in the event of mains power failure Resettable overcurrent breaker required on both live and neutral supply lines Electrical protection by resettable circuit breakers in both live and neutral supply lines Mains supply cable to be at least 3m in length Charger electrical source requirements: Amperage: ______; Voltage: ______; Frequency: ______; Phases: ______. Protections against over-voltage and over-current line conditions. |
| ACCESSORIES, CONSUMABLES, SPARE PARTS, OTHER COMPONENTS |  |  |
| 11 | Accessories (if relevant) | Breathing masks / tracheostomy connector (two sets) Filters sufficient for two weeks’ continuous use Circuit support arm. External power supply. 2 oxygen high pressure supply hose. 4 reusable adult circuits. 4 reusable or disposable pediatric circuits |
| 12 | Sterilization process for accessories (if relevant) |  |
| 13 | Consumables / reagents (if relevant) |  |
| 14 | Spare parts (if relevant) | Rechargeable batteries with at least the following characteristics: • automatic switch from AC power electric-line mode to battery operating mode and vice-versa • equipment able to operate from AC power source and external battery (12V or 24V); • Continuous monitoring working time in battery operating mode with standard ventilation not less than 5 hours; • Integrated batteries charger; • Low battery visual alarm; • 100% high capacity batteries with re-charging time not greater than 5 hours. |
| 15 | Other components (if relevant) |  |
| PACKAGING |  |  |
| 16 | Sterility status on delivery (if relevant) | N/A |
| 17 | Shelf life (if relevant) | Variable |
| 18 | Transportation and storage (if relevant) | N/A |
| 19 | Labelling (if relevant) | N/A |
| ENVIRONMENTAL REQUIREMENTS |  |  |
| 20 | Context-dependent requirements | Capable of being stored continuously in ambient temperature of 0 to 50 deg C and relative humidity of 15 to 90%. Capable of operating continuously in ambient temperature of -18 to 50 deg C and relative humidity of 15 to 90%. |
| TRAINING, INSTALLATION AND UTILISATION |  |  |
| 21 | Pre-installation requirements | Local clinical staff to affirm completion of installation |
| 22 | Requirements for commissioning | Supplier to perform installation, safety and operation checks before handover |
| 23 | Training of user/s | Training of users in operation and basic maintenance shall be provided Advanced maintenance tasks required shall be documented |
| 24 | User care | Casing to be cleanable with alcohol and chlorine wipes |
| WARRANTY AND MAINTENANCE |  |  |
| 25 | Warranty | 24 months |
| 26 | Maintenance tasks |  |
| 27 | Type of service contract |  |
| 28 | Spare parts availability post-warranty |  |
| 29 | Software / Hardware upgrade availability |  |
| DOCUMENTATION |  |  |
| 30 | Documentation requirements | User, technical and maintenance manuals to be supplied in English and/or Russian language. Certificate of calibration and inspection to be provided. List to be provided of equipment and procedures required for local calibration and routine maintenance List to be provided of important spares and accessories, with their part numbers and cost. Contact details of manufacturer, supplier and local service agent to be provided |
| DECOMMISSIONING |  |  |
| 31 | Estimated Life Span | 8 years |
| SAFETY AND STANDARDS |  |  |
| 32 | Risk Classification | Class III (EU, Japan, Canada and Australia) |
| 33 | Regulatory Approval / Certification | Must be CE approved product. |
| 34 | International standards | ISO 13485:2003 Medical devices -- Quality management systems -- Requirements for regulatory purposes ISO 14971:2007 Medical devices -- Application of risk management to medical devices IEC 60601-1:2012 Medical electrical equipment - Part 1: General requirements for basic safety and essential performance IEC 60601-1-1:2000 Medical electrical equipment - Part 1-1: General requirements for safety - Collateral standard: Safety requirements for medical electrical systems IEC 60601-1-2:2007 Medical electrical equipment - Part 1-2: General requirements for basic safety and essential performance - Collateral standard: Electromagnetic compatibility - Requirements and tests ISO 5356-1:2004 Anaesthetic and respiratory equipment -- Conical connectors -- Part 1: Cones and sockets ISO 80601-2-12:2011 Medical electrical equipment -- Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators |
| 35 | Reginal / Local Standards | EU EN 1789:2007+A2:2014 |
| 36 | Regulations |  |