## ems georgia
| technical requirment | Est. unit price (USD) | No of units requested - GEO EMS | Total cost per item - GEO EMS
Infrared thermometers | Non-contact thermometer. Measurement of body temperature. Hygienic and safe-measure without skin contact. Measurement in <1 second Infrared measurement technology Display in Celsius or Fahrenheit Free high temperature alarm setting with battery | 25 | 500 | 12500
portable transport ventilator for adults pediatric and neonatal patients, aproved for use helicopters and airplanes | transport ventilator for adults, pediatric and neonatal patients (IPPV, SIMV, CPAP, PEEP mode, pressure controllated ventilation mode) peep from 0 to no less than 20 mBar (cmH20), with battery (DC12v and AC220v charging device). Multiple flow sensor, the basic and a spare countour for adult and neonatal patients, (countour with test lung and face mask) brand new, color display, wall mounting, with transportation bag, with central and local pipe of the oxygen, with own oxygen cylinder, working temperature range at least from ‐18 c to +50 c‐ , the oxygen concentration mixed and 100% , storage battery at least 5 hours working time, charging time not more than 5 hours, must have a certificate of ISO 10651-3, CE certificate, agreeng document to EN 1789. ventilator should have certificate for using at air transport. | 25000 | 50 | 1250000
Defibrilator with patient monitoring system for adults and pediatric peds | External biphasic Defibrillator, output energy levels from 5 to not less than 270 J with adult and pediatric integrated paddles, with battery, with colour display, ECG, SpO2, P index, external pacing , battery operation time in full charge at least 300 min. ISO, CE Certificate existence, with wall mounting device. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 4000 | 350 | 1400000
ECG 12 lead | 12-lead ECG with colour screen, with ECG recording manual and automatic function, with battery, charging time no more than 4 hours. With the ability to print on paper not less than 80 mm, with ECG cable, with multiple electrodes. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 1400 | 350 | 490000
ambulance car b type |  | 60000 | 50 | 3000000
ambulance car c type (intensive care mobile unit) |  | 80000 | 30 | 2400000
portable oxygen generation system (pogs) | Equipment with the ability to quickly transport and arrange oxygen in case of emergency, in order to be used in field missions. It should be possible to supply oxygen to the patient directly from the device, as well as to pump oxygen cylinders. It should be possible to provide oxygen for the breathing ventilator with this device. | 30000 | 90 | 2700000
patient monitor for vital signs | Patient monitor, with colour display, for adults, pediatric and neonatal patients options, with wall or shelf mounting device, (ECG, NIBP, SpO2, T, HR, P), ECG chanels: I, II, III, aVR, aVL, aVF, V, Equipped with a set of electrodes, transportation battery( 12W და 220W Charging device), storage working duration at least 5 hours. ISO, CE Certificate existence, with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 1500 | 20 | 30000
portable oxygen concentrator | Concentrator with a capacity not less than 5 liters of oxygen per minute, with flow regulation, with inhalation function, oxygen production not less than 93% | 1200 | 20 | 24000
Pulse oximeters | Portable Pulse oximeter, with color screen, on display - pulse, percentage of oxygen, with battery, with hanger and protective bag | 18 | 700 | 12600
Set of Tetra standard hand Radio station |  | 1000 | 490 | 490000
Set of Tetra standard Car Radio station |  | 1400 | 380 | 532000
Tetra system dispatch control and monitoring set |  | 41000 | 43 | 1763000
Field Hospital tent sets for bioisolation patients | The tent should be at least 30 sq.m. with a floor, lighting system, heating and cooling system, if necessary with a compressor, the possibility of integration with another tent, in order to manage infectious patients, in case of need the patient should be isolated in the tent. | 70000 | 4 | 280000
isolation patient transportation capsule | Transport isolation capsules, especially for transportation of patients with dangerous infections, should be available medical manipulations with the patient placed in the capsule, if necessary, without unisolation, with a set of appropriate filters and a gas circulation device. | 6000 | 15 | 90000
Motocycle for ambulance firs responders with helmets for ambulance firs responders |  | 19800 | 50 | 990000
CPR machine lucas 2 | Automatic cardiac resuscitation device, the ability to automatically perform chest compressions without the involvement of medical personel, frequency not less than 100 pressures per minute | 15000 | 150 | 2250000
Bone Catheter for adults and pediatric patients | Bone catheter for adults and pediatric patients (1000 pieces for adults and 500 pieces for pediatric) | 85 | 1500 | 127500
Rec. camera for ambulance patient cabine | Recorder with data storage and transmission (wifi / lan). It should be possible to start / finish recording manually as well as set it to automatic mode | 300 | 400 | 120000
storage for servers |  | 62000 | 1 | 62000
server |  | 31000 | 4 | 124000
Tablet for Ambulance Emergency Medical Care Car | Tablets for ambulances, with the appropriate standard of drop, vibration safety, GPS, WIFI function, with the ability to support at least 4 G Internet, equipped with charger, car mount device | 800 | 400 | 320000
Body worn camera for ambulance group | body worn recording camera with data storage and transmission capability,with manually start / finish recording function, with the ability to store data for at least 24 hours | 560 | 100 | 56000
TOTAL COST: |  |  |  | 18523600

## consumable goods
| Est. unit price (USD) | No of units requested - GEO EMS | Total cost per item - GEO EMS
Gloves | 0.07 | 1000000 | 70000
Gowns | 0.8 | 2000000 | 1600000
Goggles | 13 | 50000 | 650000
Surgical masks | 0.66 | 3000000 | 1980000
N95 respirator masks | 0.66 | 300000 | 198000
N95 mask fit test kit | 400 | 5 | 2000
Antibacterial liquid soap (liters) | 10 | 50000 | 500000
Alcohol-based hand rub (70%, liters) | 12 | 100000 | 1200000
Chlorine-based cleaning solution for surfaces | 15 | 10000 | 150000

## Primary health care
| technical requirment | Est. unit price (USD) | No of units requested - GEO EMS | Total cost per item - GEO EMS
ECG 12 lead | 12-lead ECG with colour screen, with ECG recording manual and automatic function, with battery, charging time no more than 4 hours. With the ability to print on paper not less than 80 mm, with ECG cable, with multiple electrodes. with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 1400 | 950 | 1330000
patient monitor for vital signs | Patient monitor, with colour display, for adults, pediatric and neonatal patients options, with wall or shelf mounting device, (ECG, NIBP, SpO2, T, HR, P), ECG chanels: I, II, III, aVR, aVL, aVF, V, Equipped with a set of electrodes, transportation battery( 12W და 220W Charging device), storage working duration at least 5 hours. ISO, CE Certificate existence, with connection posibility via inetrnet or connection posibility with other relevant device, for transfering data (wifi/lan) | 1500 | 950 | 1425000
portable oxygen concentrator | Concentrator with a capacity not less than 5 liters of oxygen per minute, with flow regulation, with inhalation function, oxygen production not less than 93% | 1200 | 950 | 1140000
Pulse oximeters | Portable Pulse oximeter, with color screen, on display - pulse, percentage of oxygen, with battery, with hanger and protective bag | 18 | 1300 | 23400
TOTAL COST: |  |  |  | 3918400