SPECS for TRANSPORT VENTILATOR

The World Bank assists client (Government of Georgia) to seek critical care ventilators (emergency transport use) that meet the following technical specifications. Also indicate which accessories/supplies (e.g., breathing circuits) you can supply for one year of continuous use. If additional supplies are needed for continuous operation that you do not provide, clearly indicate this and provide appropriate distributor information, if available.

Technical Specifications

- Compliant with ISO 80601-2-80 and ISO 80601-2-79

- Model must match the voltage and frequency of the purchasing country’s local power grid (e.g., 110-120 VAC at 60 Hz or 220-240 VAC at 50 Hz)

- Tidal volume up to 1000 mL

- Pressure (inspiratory) up to 80 cm H20

- Volume (inspiratory) up to 120 L/min

- Respiratory rate (RR): up to 60 breathes per minutes (BPM)

- Synchronized intermittent mandatory ventilation (SIMV): RR up to 40 BPM

- CPAP/PEEP up to 20 cm H20

- Pressure support up to 45 cm H20

- FiO2 between 21% and 100%

- Inspiratory and expiratory times up to at least 2 sec and 8 sec respectively

- I:E ratio from 1:1 to 1:3

- Air and externally supplied oxygen mixture ratios full controllable

- Inlet gas supply (O2) pressure range 35 psi to 65 psi

- Includes integrated medical air compressor with inlet filter

Modes of Ventilation

- Volume controlled

- Pressure controlled

- Pressure support

- SIMV with pressure support

- Assist/control mode

- CPAP/PEEP

Required Alarm Conditions

If alarms can be silenced, this feature is temporary and clearly displayed on the ventilator’s screen when annunciated

- FiO2, minute volume, pressure, PEEP, apnea, occlusion, high RR, circuit disconnection

Power failure, gas disconnection, low battery, vent inoperative, self-diagnostics

Training and Manuals

- User Training:  Supplier shall provide in-person or on-line training adequate to meet the needs of all operators. Training shall be available in English, Spanish and French, and shall be available to users within two (2) weeks of equipment delivery.

- User Manuals:  Supplier shall provide at least one (1) user manual for each device purchased. Manuals shall provide guidance on the operation of all features of the purchased device, be provided in at a printed format, and be available in English, Russian.

- Service Manuals:  Supplier shall provide at least one (1) maintenance manual for each device purchased. Manuals shall provide at least, equipment schematics, parts summaries, preventive maintenance requirements, and troubleshooting guidance, be accessible in printed formats, and be available in English, Russian.