## Equipment
HOSPITAL EQUIPMENT | 55
Pulmonary Ventilator: portable (pediatric and adult) with oxygen tank different breathing regimes. | 10
basic patient monitor system: (ECG, Resp, HR, NIBP,Temp, SpO2,) with touch screen. | 45

## Diagnostics
Diagnostics | Min Specs | # of pcs
nCov-2019 Rapid detection kit (IgM/IgG) | Utilizes human whole blood (preferably peripheral), serum, or plasma· Built-in control (sensitivity) ≥ 90%, specificity ≥ 95%, number of samples used for clinical evaluation ≥ 100; · Absence of cross-reactions: human coronavirus panel, HBV, HCV HIV-1 HIV-2 Adenovirus, Parainfluenza virus 1-4, Influenza A, Influenza B, Respiratory syncytial virus, Rhinovirus· Preferably: Follow up with additional positive and negative control samples, sample collection kit and cross-examination: Human Metapneumovirus (hMPV) Chlamydia pneumoniae Streptococcus pneumoniae Mycobacterium tuberculosis Mycoplasma pneumoniae EB Virus, Enterovirus 71· Used in rapid, qualitative and differential detection of IgG and IgM antibodies· Delivers clinical results between 10 and 20 minutes at the point-of-care· Visual interpretation of results· No special equipment needed· CE-Marked, IVD, Safe Packaging, Labeling Method: Immunochromatographic· STANDARD: to be approved by FDA | 200000
nCov-2019 Rapid detection kit (Antigen) |  | 200000
nCOV-19 Real-time PCR (with extraction kits) | The kit should be multiplex assay with capability of detection of at least two target genes and internal extraction control. Negative and positive PCR controls as well as master mix should be included in the kit.The limit of detection with 95% confidence interval (LoD 95) for screening / confirmatory assays should be at least 8 / 10 copies per reaction.The specificity of the kit should be tested to demonstrate the analytical specificity to Coronavirus SARS-CoV-2. It should not show the cross amplification with other human pathogens including human coronaviruses (229 E, HKU1, MERS, SARS, NL63, OC43), Influenza, RSV and other respiratory viruses.The limit of detection with 95% confidence interval (LoD 95) for screening / confirmatory assays should be at least 8 / 10 copies per reaction.The specificity of the kit should be tested to demonstrate the analytical specificity to Coronavirus SARS-CoV-2. It should not show the cross amplification with other human pathogens including human coronaviruses (229 E, HKU1, MERS, SARS, NL63, OC43), Influenza, RSV and other respiratory viruses. | 200000

## PPE
PPE | Min. Specs | # of pcs
GOGGLE |  | 150000
GOGGLES PROTECTIVE, wraparound, soft frame, indirect vent. |  | 150000
GOWN |  | 2100000
SURGICAL GOWN, non sterile, disp., size L | length mid-calf; EU PPE Regulation 2016/425 and EU MDD Directive 93/42/EEC; FDA Class I or II medical device, or equivalent; EN 13795 any performance level, or equivalent; AAMI PB70 all levels acceptable, or equivalent | 300000
SURGICAL GOWN, non sterile, disp., size M |  | 300000
SURGICAL GOWN, non sterile, disp., size XL |  | 300000
SURGICAL GOWN, non sterile, disp., size XXL |  | 300000
ISOLATION GOWN, coat type, non sterile, disp., | Material - terylene taff + TPU impermeable membrane with a weight of more than 70 grams; PPE directive 89/686/EEC or equivalent; EN 340 level or equivalent; CE marked; FDA approved | 300000
SURGICAL GOWN, sterile, disp., size L | Laminated; MDD Council Directive 93/42/EEC or equivalent; CE Marked; FDA Aproved; Wieght not less than 45 gr | 150000
SURGICAL GOWN, sterile, disp., size M |  | 150000
SURGICAL GOWN, sterile, disp., size XL |  | 150000
SURGICAL GOWN, sterile, disp., size XXL |  | 150000
GLOVE |  | 2000000
GLOVE EXAMINATION, nitrile, pf, size L | • EU MDD Directive 93/42/EEC Category III | 500000
GLOVE EXAMINATION, nitrile, pf, size M | • EU PPE Regulation 2016/425 Category III | 500000
GLOVE EXAMINATION, nitrile, pf, size S | • EN 455 / EN 455 | 500000
GLOVE EXAMINATION, nitrile, pf, size XL | • ASTM D6319 or equivalent | 500000
MASK |  | 2500029
MASK SURGICAL, type IIR, level 2, s.u, non sterile, earloop, size L |  | 500000
MASK SURGICAL, type IIR, level 2, s.u, non sterile, earloop, size M |  | 500000
MASK SURGICAL, type IIR, level 2, s.u, non sterile, earloop, size S |  | 500000
RESPIRATOR, mask, FFP2/N95, type IIR, s.u., unvalved, noseclip |  | 1000000
N95 mask fit test kit |  | 29

## Data Validation (hidden)
Selection:
Available through existing LTA/Contracts
Available stock
Not available