| nCov-2019 Rapid detection kit (Ig M/IgG) | Utilizes human whole blood (preferably peripheral), serum, or plasma· Built-in control (sensitivity) ≥ 90%, specificity ≥ 97%, number of samples used for clinical evaluation ≥ 100; · Absence of cross-reactions: human coronavirus panel, HBV, HCV HIV-1 HIV-2 Adenovirus, Parainfluenza virus 1-4, Influenza A, Influenza B, Respiratory syncytial virus, Rhinovirus· Preferably: Follow up with additional positive and negative control samples, sample collection kit and cross-examination: Human Metapneumovirus (hMPV) Chlamydia pneumoniae Streptococcus pneumoniae Mycobacterium tuberculosis Mycoplasma pneumoniae EB Virus, Enterovirus 71· Used in rapid, qualitative and differential detection of IgG and IgM antibodies· Delivers clinical results between 10 and 20 minutes at the point-of-care· Visual interpretation of results· No special equipment needed· CE-Marked, IVD, Safe Packaging, Labeling Method: Immunochromatographic· STANDARD: to be approved by FDA |
| --- | --- |
| nCov-2019 Rapid detection kit (Antigen) | COVID-19 Ag test based on the rapid chromatographic immunoassay for the qualitative detection of specific antigens to SARS-CoV-2 present in human nasopharynx. Specimen type - Nasopharyngeal swab. Delivers result no later than 30 mins. Easy to use (no more than four steps). Included internal control - procedural control detecting the capability of the assay. Desired if pack includes all accessories needed for taking sample and its application to test & no equipment needed. Clinical sensitivity greater than 84% (within 95% confidence intervals), Clinical specificity greater than 98% (within 95% confidence intervals). Absence of cross-reactivity with other human coronaviruses. Regulatory status – CE-marked, IVD, preferable approved by FDA. Reagent storage temperature 2- 30°C. |